## Slide 1
Development of a
Scintillation Fiber Transverse Profile Monitor for Low-Intensity Ion Beams at HIT
14.07.2022, Richard Hermann

## Slide 2: Agenda
Intro
Prototype
Experiments
Conclusion & Outlook

## Slide 3: DFG Project “Energy-painted ion radiography for precision radiotherapy“
Helium Beam Radiography (Tim Gehrke)
  Smaller geometric uncertainties between CT and therapy.
  Tissue stopping power more precise.
  Smaller radiation dose.
Requirement:
Monitor for low intensity ion beams. Beam pos. and shape not yet monitored for ≤ 10^5 pps (feedback missing)
*G. Aricò, T. Gehrke, J. Jakubek, R. Gallas, S. Berke, O. Jäkel, A. Mairani, A. Ferrari, M. Martišíková, Investigation of mixed ion fields in the forward direction for 220.5 MeV/u helium ion beams: comparison between water and PMMA targets, Physics in medicine and biology 62, 8003 (2017).
*
100 Ionen

## Slide 4: Idea from LHCb Tracker @ Cern
LHCb fibers: 250µm x 3m, 6 layer, round (Kuraray SCSF-78MJ) (B. Leverington)
7.200 photons per MeV
  0,4 MeV/mm (@220 MeV/u He)
  5,4% trapping
  35% PDE* der SiPM**
  2x ~5% opt. coupling losses (air) (Rs = 11.55% , Rp = 1.19%)
  ≈ 45 photo electrons / He-ion (> 5 p.e. needed for S/N threshold)
  + 5% non-detected ions in on layer of round fibers.
*PDE: Photon Detection Efficiency **SiPM : Silicon Photomultiplier
B. Leverington „The SciFi Tracker for the LHCb Upgrade“, Vortrag 2014 in Neckarzimmern
Notes: What?
Lhcb experiment at cern
They have small fibres, which glow when a ion is passint through – these are the szintillating fibres
Here you see that they are actually quite big.
With photomultipliers they amplify the light-signal
to get a sense of the scale we’re talking about: in my setup I get about 50 photons per Helium-Ion – and I don’t know yet how many I loose due to bad optical coupling between fibre and photomultiplier.

## Slide 5: Profile monitor for low intensity ion beams Advantages of LHCb inspired SciFi-Detector
Highly-tested, radiation hard and cheap fibers
  Low running/replacement costs: ion damage and aging only on fibers
  Expensive electronics in safe distance to beam (≠ semiconductor det.)
Clear structure of detector system
  Only a few components
  No gas or vacuum necessary
  No semiconductor creation necessary
  No cooling necessary (works at RT)
Usability
  Scalability of detector channels and electronics
  Readout electronics commercially available and comes with control software for first experiments

## Slide 6: Prototype Idea for first test setup
Readout inspired by IDEA Dual-Readout calorimeter - R. Santoro, INFN

## Slide 7: Scintillating fibers Probes
Cutted and polished by hand
PI Heidelberg: gluing into 3D print and milling
Kuraray SCSF-3HF multiclad
  Radiation harder
  7ns decay, 530nm
Kuraray SCSF-78 multiclad
  Same as LHCb
  2.8ns decay, 450nm
Radiation hardness: SCSF-3HF 20% loss where SCSF-78 has 60% loss *
Kuraray 3HF multi clad
Kuraray 78 multi clad
* Preliminary radiation damage analysis for the HIT beam profile monitor; M.Dziewiecki, B.Leverington, G.Meo; Heidelberg 13.02.2018

## Slide 8: Photomultiplier SiPMs on PCB*
MPPC** (S13360-1350PE)
  667 cells (50µm APDs) on 1,3x1,3mm² active area
  Peak PDE at 450nm = scintillating fiber (SCSF-78)
*PCB: Printed Ciruit Board
**MPPC: Multi-Pixel Photon Counter (by Hamamatsu)
≈ 9 cm

## Slide 9: Electronics Readout System
CAEN FERS A5202 [64 Channel CITIROC ASIC]
Future: DT5215 (CAEN) combines several FERS boards
User Manual UM7945, A5202/DT5202; 64-Channel Citiroc-1A Unit for FERS-5200; Rev. 2 - September 9th, 2021

## Slide 10: Prototype Setup Realization
Scintillating Fibres
SiPMs on PCB
CAEN FERS A5202
Ion Beam

## Slide 11: Measurements In Beam 1H, E#255, F#4, < I#1
≈ 100.000 ions/sec.
Threshold 350 (a.u.) from staircase plot
Signal to DCR: ≈ 20.000
Signal to Noise within beam: ≈ 50 - 100
logarithmic

## Slide 12: Proof of Principle: very low intensity beam 1H, E#255, F#4, << I#1
Really low intensity: ≈ 8 ions/sec. (-> from TimePIX)
Edge of useful resolution reached.
Proof of principle
✓
Simulation: 100 Ions
Notes: Only some of the spills have visible signal – some actually seem empty

## Slide 13: Measurements In Beam 1H, E#255, F#4, < I#1
Kuraray SCSF-78 (450nm, 8k phot./MeV*)
VS
Kuraray SCSF-3HF (530nm, 7.1k phot./MeV*)
I
Fibre-SiPM combination seem to have better S/N in this application.
SiPM PDE
≈ 30.000 ions/sec .
≈ 60 ions/sec .
* From Saint-Gobain fibre product sheet. Values for Minimum Ionizing Particles (MIP), corrected for PMT sensitivity

## Slide 14: Measurements In Beam C12 Scan, 430 MeV/u, F 3.4mm, 20k ions/mm
100 steps in 1,2 sec (from -50mm till 50mm)  2E6 ions/sec (≙ I#1)
Position and width, movement and timing measurable.
If perfected, this can be used to define supply voltage corrections for SiPMs
10ms frames
8x slower than reality
Notes: If not working : time is coded in colour on left side
More visual on right side

## Slide 15: Ion Comparison All at E#10 and I#1
I = 10mA was max.
I > 1mA looses gaussian shape
Over current only for wide (F > 1), high intensity proton beams.
Anyway, data loss through jumped time windows
Over a certain limit (roughly at 1e7 pps) SiPMs can’t fully recharge before next hit.  Missing counts  non-gaussian form

## Slide 16: Energy comparison Oxygen with E#10 and E#255 (1.5e6 pps)
Higher Energy (same intensity)  smaller spot (=FWHM)
Area A is the same difference: 0,01%
y0 = x ± ~8%
xc = x ± ~0,1%
w = x ± ~0,13%
A = x ± ~0,12%
W, FWHM, σ: Faktor /2,34
Height: Faktor *2,34
Intensity determinable by area A of the fit.

## Slide 17: Timing Mode Angle Comparison
C12 E255 F4 I1 -91.??°

## Slide 18: Timing Mode Comparison with integrated scintillator
“Run”: my detector
“H1DI1P…”: integrated scintillator in accelerator
 Spill form is really good comparable (except intensity fluctuations)
C12 E255 F4 I1 -91.3°

## Slide 19: Conclusion
Prototype successfully proven to measure:
  Beam position and width
  Intensity (spill form)
  Tracking single ions (time stamp)
Limits:
  1.8 MB data throughput, due to “how data is packed and transmitted in the links“
    one full set of data takes 2776 bits (≈ 350 bytes)
  Counting: Data loss, if too short time intervals, e.g <200µs resolution. Data loss, if to many ions/s, e.g >10^7 ions/s.
  Timing: Data loss, if to many ions/s, e.g >5*10^4 ions/s.

## Slide 20: Outlook
More and more direct comparison to BAMS (HIT), Timepix (Tim) and SciFi-monitor (PI).
Counting mode will be faster with Zero-Suppression (Caen promised program update within weeks)
FERS A5202 & DT5215 for second detection layer are in order process and will be delivered in December & January
  Concentrator Board and TDlink will allow data throughput of 60 MB/s (~170 000 events/s possible: 200µs -> ~6 µs time windows in counting)
  Prototype extension: vertical plane & quadratic fibers.

## Slide 21: End Questions?
With special thanks to:
HIT
Andreas Peters
Prof. Thomas Haberer (Supv.)
Michael Galonska
Christian Schömers
Eike Feldmeier
Andreas Gaffron
Stefan Brons
Rosi Vay-Meyer
DKFZ
Tim Gehrke
Maria Martisikova
Phys. Inst. HD
Blake Leverington
Sebastian Bachmann
Uni Frankfurt
Prof. Ulrich Ratzinger (Supv.)
UK HD
Prof. Oliver Jäkel
INFN
Romualdo Santoro

## Slide 22: Appendix Extra Slides

## Slide 23: Dark Count Rate (DCR) measurement Threshold Scan
Different Threshold-Values: 220, 300, 340, 360 (arbitrary units)
 Threshold um 350 seems good
1 p.e. are due to dark counts
2 p.e., 3 p.e., etc. are due to crosstalk
S-curve / staircase plot
https://hub.hamamatsu.com/jp/en/technical-note/how-sipm-works/index.html
Notes: Channels equals millimeters as I use 1mm fibres.
So 64 channels are 64 mm (with 1mm resolution) in y-direction
Counts per second means: the summed clics (fire processes) of a SiPMs – every channel is a SiPM

## Slide 24: State of the Art INFN – R.Santoro – INFN
IDEA Dual-Readout
calorimeter collaboration
- R. Santoro (INFN)
Also:
SiPM + FERS A5202
Romualdo Santoro, „IDEA DR Calorimeter: readout electronics“ , Univeristàdell’Insubria and INFN-MI, 2021

## Slide 25: Dark Count Rate (DCR) measurement Strange peaks
Light leaks from behind through pcb through-holes.
Status-LEDs

## Slide 26: Dark Count Rate (DCR) measurement Strange peaks
Light leak through through-holes
And: Gap on the side
   3D print for light protection
tight
gap

## Slide 27: I Szintillationsfasern

## Slide 28: Scintillator Organic
https://scionix.nl/frame/
E. Auffray, „Scintillation Detectors“, CERN, EP-CMX technoweek School 2021
Scintillation: Alternation of free valence electrons between molecular orbitals. Typical: Benzene ring circumferential e-
Pro: Fast (ns) - intensity determinable, magnetic field insensitive, fibres enable position measurements, (+ advantages slide 6)
Contra: No energy measurement possible (non-linearity & saturation), defined radiation hardness and therefore defined lifetime – can’t be refurbished
Notes: S: Elekton energy levels (excited states) – within these levels are molecular vibration modes (fs decay)
S = Total Spin Quantum Number  Singulet: antiparallel 2*S+1= 1 / Triplett parallel spin 2*S+1=3
pi-Elektronen: nicht an ein atom gebunden, sondern nicht lokalisierbar im benzol-ring „Molekülorbital“
Transitions: allowed, forbidden, non-radiant, …
Szintillating process has something to do with light generation out of ions passing through material.
Anorganic scintillators: crystals for example used in PET scanner.
Organics one could combine with plastic, or are itself some kind of plastic.
Fast: Excitation & then Nanoseconds decay via flurescense

## Slide 29: Scintillator Plastic scintillators and fibers
E. Auffray, „Scintillation Detectors“, CERN, EP-CMX technoweek School 2021
1. Primary ionization => excitation of molecules in the base polymer
2. De-excitation of the base polymer [theoretically produces scintillation photons (~ 300 nm)] => Energy transfer directly to the fluor in a very short distance.
3. Primary fluor emits at a longer wavelength (~ 340 nm),
4. Absorbed by a secondary fluor
5. Secondary fluor emits in the visible (~ 400 nm)
6. Detect by photodetector
Stokes-Shift: Difference between maxima in absorption and emission.
  Carrier medium
  Primary scintillator
  Wavelength shifter
Notes: - mechanism of energy absorption: energy is first absorbed by the solvent, then passed onto the scintillation solute (the details of the transfer are not clearly understood)
- Emission has longer wavelength (lower energy)

## Slide 30: Scintillating fibers Radiation hardness
Kuraray SCSF-3HF multiclad
  3-hydroxyflavone
  Smaller radiation damage than other fibers *
  Lifetime & attenuation length also dependent on bending and on temperature of the fiber. *
  annealing effect in days/weeks/months, total and none at all **
  20% loss where Kuraray SCSF-78 multiclad has 60% loss ***
Kuraray 3HF multi clad
Kuraray 78 multi clad
*Radiation damage in scintillating fibers -Preliminary literature study- ; Michael Moll, CERN, 15.10.2003
** A fiber detector radiation hardness test, J Bähr et al., Nuclear Instruments and Methods in Physics Research Vol 449, 11.08.2000
*** Preliminary radiation damage analysis for the HIT beam profile monitor; M.Dziewiecki, B.Leverington, G.Meo; Heidelberg 13.02.2018

## Slide 31: Scintillating fibers Cladding
E.g. Kuraray “polystyrene core (inner side) with a fluorescent agent and a methacrylate cladding (outer side)”
https://www.kuraray.com/products/psf
Notes: Deshalb trapping recht schlecht

## Slide 32: II SiPM

## Slide 33: Photomultiplier Silicon Photomultiplier (SiPMs)
APDs: 50 x 50 µm² Avalanche-Photodiode: Semiconductor equivalent of a photomultiplier.
SiPM: 50µm APDs on 1,3x1,3mm² active area = 667 pixels
Pro: Fast (ns) - intensity determinable, magnetic field insensitive; (+ advantages slide 6)
Contra: No energy measurement possible (Geiger-Mode), radiation damage – aging
**Hamamatsu, MPPC Technical Note Cat. No. KAPD9008E01 Apr. 2021 DN
**
*Von Kirnehkrib - Eigenes Werk, CC BY-SA 3.0, https://commons.wikimedia.org/w/index.php?curid=17725203
*
Notes: SiPMs consist of …

## Slide 34: Detector principal SiPM (Silicon-Photomultiplier)
The primary waveforms (1 p.e.) are due to dark counts; whereas, 2 p.e., 3 p.e., etc. are due to crosstalk
https://hub.hamamatsu.com/jp/en/technical-note/how-sipm-works/index.html

## Slide 35: III FERS

## Slide 36: Electronics The CAEN FERS
Measurement Modes:
  Counting Mode no dead time (except saturation), max. 20Mcps/channel
  Timing Mode TOA (25-bit), dynamic range ~16,78ms (2^25 *0.5ns)
  Time Stamped Spectroscopy: TOA (16-bit) + TOT (9-bit), dynamic range ~32,77µs
  Spectroscopy Mode Amplitude conversion by 13-bit ADC (~10µs dead time)
Data Throughput: 1.8 MB/s (due to data packaging  200µs time windows)
User Manual UM7945, A5202/DT5202; 64-Channel Citiroc-1A Unit for FERS-5200; Rev. 2 - September 9th, 2021

## Slide 37: Readout FERS A5202 (CAEN)
User Manual UM7945, A5202/DT5202; 64-Channel Citiroc-1A Unit for FERS-5200; Rev. 2 - September 9th, 2021

## Slide 38: Readout FERS A5202 (CAEN)
User Manual UM7945, A5202/DT5202; 64-Channel Citiroc-1A Unit for FERS-5200; Rev. 2 - September 9th, 2021

## Slide 39: Readout FERS A5202 (CAEN)
User Manual UM7945, A5202/DT5202; 64-Channel Citiroc-1A Unit for FERS-5200; Rev. 2 - September 9th, 2021

## Slide 40: IV Simulation

## Slide 41: Sideproject – Simulation Geant4 LHCb-tracker Simulation adapted
Only first adaptions: size, material, readout cells
6x250µm fiber mat 1x1mm fiber mat
BUT Output file/-plots not yet useable (e.g. pixels not correctly assigned)
Kein Signal
Crosstalk